# BlindVision — Supplementary Material

Supplementary Tables S1–S18 accompany the main manuscript. They contain protocol, dataset, failure-mode, power, reliability, software-environment, computational, comparison, navigation, ablation, sensor, latency, subsystem-result, detection-metric, and reconciliation details moved out of the main manuscript to improve readability.

TABLE S1. Dataset composition.

| Split | Images | Annotated objects |
| --- | --- | --- |
| Train | 5,600 | — (not reported per-split) |
| Validation | 700 | — (not reported per-split) |
| Test (held-out) | 700 | 842 |

TABLE S2. Detection success by lighting and occlusion.

| Condition | Detection success rate |
| --- | --- |
| Bright daylight | 97.8% |
| Low light | 87.9% |
| No occlusion | 97.5% |
| Severe occlusion (>75%) | 74.2% |

TABLE S3. Experimental protocol summary.

| Item | Value |
| --- | --- |
| Evaluation sessions (full prototype) | 20 |
| Total navigation trials | 100 |
| Approx. total walking distance | ~2 km |
| Object detection attempts | 200 |
| BLE packets transmitted (test) | 8,000 |
| Visually impaired (BVI) participants, total | 7 (1 earlier + 6 later) |
| Additional BVI-specific trials (six-participant session) | 48 (8 per participant: 3 outdoor nav, 3 stair/drop-off, 2 water hazard) |
| Sighted testers | Project-team members (count not specified) |

TABLE S4. Power budget.

| System | Approx. input power / basis |
| --- | --- |
| Smart Goggles (5-V USB input) | 12.11 W mean, measured from five released active-detection samples |
| Smart Stick (3.7-V nominal cell) | ≈0.8 W, derived/assumed for runtime projection |
| Combined | Not directly measured; separate power domains |

TABLE S5. Battery/runtime characterization.

| System | Measured / derived runtime | Basis |
| --- | --- | --- |
| Goggles — idle | Not reported | No idle samples in released power log |
| Goggles — active detection | ≈3.1 h nominal-equivalent | 37 Wh nominal / 12.11 W measured mean; projection, not depletion test |
| Goggles — maximum simultaneous load | Not reported | No maximum-load samples in released power log |
| Smart Stick | ≈8.7 h upper bound | Derived from 5,000 mAh nominal capacity at 575 mA assumed average draw; no continuous discharge test |

TABLE S6. Reliability (field measurements).

| Metric | Result |
| --- | --- |
| Continuous operating duration tested | 4 hours |
| Crashes observed | 0 |
| Restarts observed | 0 |
| Sustained RAM growth observed | 0 (none) |

TABLE S7. Software environment.

| Component | Version / detail |
| --- | --- |
| On-device detection model | YOLOv8n (nano), CPU-only, no hardware accelerator |
| Host (full prototype) | Raspberry Pi 4 Model B, 4 GB |
| Operating system | Raspberry Pi OS (exact build/architecture not specified in the release) |
| ultralytics (Smart Goggles) | >=8.2.0 |
| ultralytics (proof-of-concept cloud server) | ==8.2.28 |
| fastapi | ==0.111.0 |
| uvicorn[standard] | ==0.30.1 |
| pydantic | ==2.7.4 |
| opencv-python | >=4.9.0 |
| bleak (BLE) | >=0.22.0 |
| pyttsx3 (offline TTS) | >=2.90 |
| pyserial | >=3.5 |
| pynmea2 (GPS parsing) | >=1.19.0 |
| numpy | >=1.26.0 |
| pillow (proof-of-concept) | ==10.3.0 |
| python-multipart (proof-of-concept) | ==0.0.9 |

TABLE S8. Ten-class confusion matrix on the 700-image held-out test split.

| True \ Predicted | Person | Door | Chair | Backpack | Laptop | Bottle | Pole | Vehicle | Bicycle | Stairs |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Person | 110 | 0 | 1 | 0 | 0 | 0 | 1 | 2 | 0 | 1 |
| Door | 1 | 75 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 |
| Chair | 0 | 0 | 88 | 1 | 0 | 0 | 0 | 0 | 0 | 3 |
| Backpack | 0 | 1 | 0 | 61 | 0 | 1 | 0 | 0 | 1 | 1 |
| Laptop | 0 | 0 | 0 | 0 | 53 | 1 | 0 | 0 | 0 | 1 |
| Bottle | 0 | 0 | 0 | 0 | 0 | 66 | 1 | 0 | 0 | 3 |
| Pole | 1 | 0 | 0 | 1 | 0 | 0 | 120 | 0 | 0 | 2 |
| Vehicle | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 84 | 1 | 1 |
| Bicycle | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 57 | 1 |
| Stairs | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 0 | 91 |

TABLE S9. Computational-load measurements from the released run.

| Metric | Condition | Measured value |
| --- | --- | --- |
| CPU utilization (avg., 4 cores) | During active detection + fusion | 61% (psutil) |
| CPU utilization (peak) | During active detection + fusion | 87% (psutil) |
| RAM usage | During active detection + fusion | 1.7 GB (psutil) |
| Processor temperature | During active detection + fusion | 61 °C (vcgencmd) |
| Storage utilization | 64 GB microSDXC (OS, models, logs) | 12.3 GB occupied / 51.7 GB free |
| Camera pipeline frame rate (FPS) | Four concurrent streams, 60 s run (Section VI.I) | 22.8 FPS aggregate; 5.7 FPS per-camera mean (measured) |
| Detection stage speed (single stream) | Per frame, on-device (Raspberry Pi 4) | 122 ms/frame (from Table S15) |
| Detection stage speed (four concurrent streams) | Per camera, four streams active | 175 ms/frame (derived as 1/5.7 FPS) |
| Parallel efficiency, four streams | 22.8 FPS measured vs 32.8 FPS if perfectly parallel | 69.5% (derived) |

TABLE S10. Contextual comparison with related assistive-navigation systems.

| System | Vision | Ultrasonic/ground | Offline fallback | Caregiver layer | Latency | End-to-End Detection/Guidance Success |
| --- | --- | --- | --- | --- | --- | --- |
| ALVU [2] | No (ToF array, not camera-based) | Yes (waist-level ToF) | N/R | No | N/R | Qualitative, 162 trials |
| Gao et al. [3] | Yes (RGB+D glasses) | No | No (phone-dependent) | No | <320 ms | 100% avoidance |
| 7Sense SuperBrain 1 [6] | Yes (RGB/depth/thermal) | No | N/R | No | N/R | N/R (commercial; no published trial data) |
| Mai et al. [10] | Yes (RGB-D camera) | Yes (2D LiDAR on cane, not ultrasonic) | N/R (single integrated device, not two independently-operable nodes) | No | N/R | 61.2% overall recognition rate across 86 object classes, 25–26 FPS (as reported in [10]; not a directly comparable protocol) |
| BlindVision | Yes (4-camera YOLOv8) | Yes (5x ultrasonic + 2x IR) | Yes (Vision-Only/Stick-Only modes, Section III) | Backend/API event layer; no released smartphone app | 205 ms serialized stage-sum | 96.0% end-to-end detection & guidance success; 96.0% fused navigation |

TABLE S11. Navigation trial results by scenario.

| Scenario | Trials | Successful | 95% CI |
| --- | --- | --- | --- |
| Corridor navigation | 30 | 29 | 83.3–99.4% |
| Outdoor walk | 25 | 24 | 80.5–99.3% |
| Stair avoidance | 20 | 19 | 76.4–99.1% |
| Water hazard | 10 | 10 | 72.2–100% |
| Moving person detection | 15 | 14 | 70.2–98.8% |

TABLE S12. Subsystem ablation and power-basis status.

| System | Approx. input power / basis | Successful | Navigation success rate | Navigation success rate |
| --- | --- | --- | --- | --- |
| Smart Goggles (5-V USB input) | 12.11 W mean, measured from five released active-detection samples | Not reported | Not reported |  |
| Smart Stick (3.7-V nominal cell) | ≈0.8 W, derived/assumed for runtime projection | Not reported | Not reported |  |
| Combined | Not directly measured; separate power domains | 96 | 96.0% (overall navigation result) |  |

TABLE S13. Sensor reading accuracy, full system.

| Sensor | Tested range | Mean absolute error | SD / basis |
| --- | --- | --- | --- |
| Ultrasonic (historical aggregate) | 20–300 cm | 2.4 cm | 0.7 cm (n=10/point) |
| IR sensors | 5–80 cm | 1.8 cm | 0.5 cm |
| Water sensor | Wet/dry | 98.4% accuracy (reported aggregate) | n/a (binary) |
| IMU tilt | ±90° | ±2° | 0.6° |
| FSR | Light/medium/heavy pressure | 48/50 correct (96%) | 95% CI 86.5–98.9% |

TABLE S14. Reported aggregate water- and stair-hazard detection metrics.

| Metric | Value |
| --- | --- |
| Water Detection – Accuracy | 98.4% |
| Water Detection – Precision | 98.1% |
| Water Detection – Recall | 98.8% |
| Stair Detection – Precision | 96.2% |
| Stair Detection – Recall | 95.5% |
| Stair Detection – F1 | 95.8% |

TABLE S15. Single-stream processing-path stage breakdown.

| Process stage / characterization | Time (mean) | Measurement basis / dispersion |
| --- | --- | --- |
| Detection (image capture + YOLO inference) | 122 ms | N/A (not reported) |
| BLE transmission | 17.5 ms | N/A (not reported) |
| Sensor fusion | 9 ms | N/A (not reported) |
| Serialized stage-sum (not direct end-to-end) | 205 ms | Derived from stage means |
| Total alert time | 205 ms | Derived from stage means |

TABLE S16. Full-system experimental results by subsystem, with 95% Wilson score confidence intervals.

| Test | Total | Success | Rate | 95% CI |
| --- | --- | --- | --- | --- |
| Object detection | 200 | 192 | 96.0% | 92.3–98.0% |
| Voice guidance | 200 | 194 | 97.0% | 93.6–98.6% |
| BLE communication | 8000 | 7920 | 99.0% | 98.8–99.2% |
| Ultrasonic detection | 180 | 174 | 96.7% | 92.9–98.5% |
| SOS alert | 20 | 20 | 100% | 83.9–100% |

TABLE S17. Detection results and reconciled metrics for the 700-image held-out split. Precision and F1 are reconstructed from aggregate confusion/false-positive counts; mAP values remain reported outputs because raw prediction-level outputs are unavailable.

| Metric | Result |
| --- | --- |
| Precision | 92.2% (reconciled) |
| Recall | 95.6% |
| F1-score | 93.9% (reconciled) |
| mAP@0.50 | 96.8% |
| mAP@0.50:0.95 | 76.5% |
| IoU threshold | 0.50 |
| Average confidence threshold | 0.45 |
| Test images | 700 |
| Test object classes | 10 |
| Capture-to-alert latency | ~0.7–0.9 s (midpoint ~0.8 s) |

TABLE S18. Reconciled detection false-positive accounting.

| Metric | Value |
| --- | --- |
| Sensitivity | 95.6% |
| False-alarm rate (1 − precision) | 7.8% (from reconciled precision) |
| False Negative Rate | 4.4% |
| False positives per image | 0.097 (68 / 700) |

ADDITIONAL EXPERIMENTAL RECORDS

The following tables document the additional author-supplied experiments added after the earlier release. Raw row-level records are retained in data/raw/ and are not replaced by these summaries.

TABLE S19. Recomputed per-transducer calibration.

| Sensor | m | b | R² | Corrected MAE (cm) | Corrected SD (cm) | Max ¦error¦ (cm) |
| --- | --- | --- | --- | --- | --- | --- |
| U1 | 0.99869704 | 0.51718625 | 0.999894 | 0.707 | 0.940 | 1.976 |
| U2 | 1.00029776 | -0.01266986 | 0.999968 | 0.431 | 0.514 | 0.923 |
| U3 | 0.99884492 | 0.30787370 | 0.999931 | 0.583 | 0.759 | 1.540 |
| U4 | 0.99796163 | 0.40065947 | 0.999944 | 0.552 | 0.684 | 1.294 |
| U5 | 0.99962230 | 0.22446043 | 0.999984 | 0.303 | 0.367 | 0.689 |

TABLE S20. Incidence-angle characterization at 100 cm.

| Angle (deg) | n | Mean measured (cm) | Mean ¦error¦ (cm) | Mean relative amplitude |
| --- | --- | --- | --- | --- |
| -60 | 5 | 105.44 | 5.44 | 0.346 |
| -45 | 5 | 102.14 | 2.14 | 0.646 |
| -30 | 5 | 101.22 | 1.22 | 0.846 |
| -15 | 5 | 100.54 | 0.54 | 0.952 |
| 0 | 5 | 100.30 | 0.30 | 1.000 |
| 15 | 5 | 100.64 | 0.64 | 0.960 |
| 30 | 5 | 101.40 | 1.40 | 0.836 |
| 45 | 5 | 102.82 | 2.82 | 0.620 |
| 60 | 5 | 106.14 | 6.14 | 0.320 |

TABLE S21. Target-material characterization at 100 cm.

| Material | n | Mean measured (cm) | Mean ¦error¦ (cm) | SD (cm) |
| --- | --- | --- | --- | --- |
| flat_board | 5 | 101.94 | 1.94 | 0.34 |
| concrete | 5 | 101.42 | 1.42 | 0.26 |
| metal | 5 | 100.44 | 0.44 | 0.23 |
| fabric | 5 | 109.12 | 9.12 | 0.61 |

TABLE S22. Simultaneous-firing cross-talk check.

| Sensor | Solo mean at 100 cm | Simultaneous reading | Shift (cm) |
| --- | --- | --- | --- |
| U1 | 100.70 | 101.50 | 0.80 |
| U2 | 100.20 | 102.10 | 1.90 |
| U3 | 100.50 | 101.80 | 1.30 |
| U4 | 100.70 | 102.40 | 1.70 |
| U5 | 100.55 | 101.90 | 1.35 |

TABLE S23. Fusion-weight sensitivity summary.

| Configuration | Changed decisions (of 10) | Interpretation |
| --- | --- | --- |
| Baseline | 0 | Reference configuration |
| C +20% | 1 | One decision changed |
| C -20% | 0 | All preserved |
| P +20% | 0 | All preserved |
| P -20% | 1 | One decision changed |
| U +20% | 0 | All preserved |
| U -20% | 0 | All preserved |
| W +20% | 0 | All preserved |
| W -20% | 0 | All preserved |

TABLE S24. Trial-level paired subsystem ablation.

| Condition | Trials | Successes | Collisions | Missed hazards | Mean completion time (s) |
| --- | --- | --- | --- | --- | --- |
| full | 4 | 4 | 0 | 0 | 53.58 |
| no_vision | 4 | 0 | 4 | 0 | 35.90 |
| no_ultrasonic | 4 | 2 | 2 | 3 | 58.23 |
| no_fusion | 4 | 0 | 4 | 3 | 43.60 |

TABLE S25. Continuous Smart Stick discharge run.

| Elapsed (min) | Voltage (V) | Current (A) | State |
| --- | --- | --- | --- |
| 0 | 4.2 | 0.42 | active |
| 30 | 4.05 | 0.41 | active |
| 60 | 3.91 | 0.4 | active |
| 90 | 3.78 | 0.4 | active |
| 120 | 3.65 | 0.39 | active |
| 150 | 3.52 | 0.39 | active |
| 180 | 3.38 | 0.38 | active |
| 210 | 3.2 | 0.37 | active |
| 240 | 3.01 | 0.35 | cutoff_reached |

TABLE S26. Raw experimental records included in the repository.

| Raw file | Rows | Purpose |
| --- | --- | --- |
| data/raw/Ultrasonic_Calibration_Raw.csv | 80 | Per-transducer distance calibration |
| data/raw/Ultrasonic_Calibration_With_Correction_Raw.csv | 80 | Raw readings plus inverse-corrected values |
| data/raw/Ultrasonic_Angle_Sweep_Raw.csv | 45 | Incidence-angle response, five sensors × nine angles |
| data/raw/Ultrasonic_Material_Raw.csv | 20 | Five sensors × four target materials × one supplied observation per sensor/material at 100 cm |
| data/raw/Ultrasonic_Crosstalk_Raw.csv | 5 | One simultaneous-firing observation per sensor |
| data/raw/Fusion_Weight_Sensitivity_10cases.csv | 10 | Inputs for ten recorded fusion cases |
| data/raw/Fusion_Weight_Sensitivity_Summary.csv | 9 | Nine tested weight configurations |
| data/raw/Paired_Subsystem_Ablation_Raw.csv | 16 | Four paired navigation trials × four conditions |
| data/raw/Smart_Stick_Continuous_Discharge_Raw.csv | 9 | Single 240-min discharge record |

APPENDIX A. RAW RECORDS

The following tables reproduce the newly added raw author-supplied records row-for-row. They are included here in addition to the CSV files in the repository; no earlier aggregate records are replaced or deleted.

Ultrasonic_Calibration_Raw.csv

| sensor_id | true_distance_cm | measured_distance_cm | trial |
| --- | --- | --- | --- |
| U1 | 40 | 40.5 | 1 |
| U1 | 40 | 40.8 | 2 |
| U1 | 60 | 60.2 | 1 |
| U1 | 60 | 60.6 | 2 |
| U1 | 80 | 79.9 | 1 |
| U1 | 80 | 80.4 | 2 |
| U1 | 100 | 100.3 | 1 |
| U1 | 100 | 101.1 | 2 |
| U1 | 150 | 149.5 | 1 |
| U1 | 150 | 150.8 | 2 |
| U1 | 200 | 199.2 | 1 |
| U1 | 200 | 201.0 | 2 |
| U1 | 250 | 248.9 | 1 |
| U1 | 250 | 251.4 | 2 |
| U1 | 300 | 298.5 | 1 |
| U1 | 300 | 302.1 | 2 |
| U2 | 40 | 39.8 | 1 |
| U2 | 40 | 40.3 | 2 |
| U2 | 60 | 59.7 | 1 |
| U2 | 60 | 60.4 | 2 |
| U2 | 80 | 80.1 | 1 |
| U2 | 80 | 79.6 | 2 |
| U2 | 100 | 99.9 | 1 |
| U2 | 100 | 100.5 | 2 |
| U2 | 150 | 150.2 | 1 |
| U2 | 150 | 149.7 | 2 |
| U2 | 200 | 200.4 | 1 |
| U2 | 200 | 199.5 | 2 |
| U2 | 250 | 249.2 | 1 |
| U2 | 250 | 250.8 | 2 |
| U2 | 300 | 301.0 | 1 |
| U2 | 300 | 299.4 | 2 |
| U3 | 40 | 40.1 | 1 |
| U3 | 40 | 39.9 | 2 |
| U3 | 60 | 60.5 | 1 |
| U3 | 60 | 59.8 | 2 |
| U3 | 80 | 80.2 | 1 |
| U3 | 80 | 80.7 | 2 |
| U3 | 100 | 100.1 | 1 |
| U3 | 100 | 100.9 | 2 |
| U3 | 150 | 149.8 | 1 |
| U3 | 150 | 150.4 | 2 |
| U3 | 200 | 201.2 | 1 |
| U3 | 200 | 198.9 | 2 |
| U3 | 250 | 250.3 | 1 |
| U3 | 250 | 249.1 | 2 |
| U3 | 300 | 298.8 | 1 |
| U3 | 300 | 301.5 | 2 |
| U4 | 40 | 40.4 | 1 |
| U4 | 40 | 39.6 | 2 |
| U4 | 60 | 60.1 | 1 |
| U4 | 60 | 60.7 | 2 |
| U4 | 80 | 79.8 | 1 |
| U4 | 80 | 80.3 | 2 |
| U4 | 100 | 100.4 | 1 |
| U4 | 100 | 101.0 | 2 |
| U4 | 150 | 150.6 | 1 |
| U4 | 150 | 149.2 | 2 |
| U4 | 200 | 199.8 | 1 |
| U4 | 200 | 200.7 | 2 |
| U4 | 250 | 251.1 | 1 |
| U4 | 250 | 248.6 | 2 |
| U4 | 300 | 300.2 | 1 |
| U4 | 300 | 299.1 | 2 |
| U5 | 40 | 39.9 | 1 |
| U5 | 40 | 40.2 | 2 |
| U5 | 60 | 60.3 | 1 |
| U5 | 60 | 59.9 | 2 |
| U5 | 80 | 80.0 | 1 |
| U5 | 80 | 80.5 | 2 |
| U5 | 100 | 100.3 | 1 |
| U5 | 100 | 100.8 | 2 |
| U5 | 150 | 149.9 | 1 |
| U5 | 150 | 150.5 | 2 |
| U5 | 200 | 200.1 | 1 |
| U5 | 200 | 199.8 | 2 |
| U5 | 250 | 249.7 | 1 |
| U5 | 250 | 250.4 | 2 |
| U5 | 300 | 299.6 | 1 |
| U5 | 300 | 300.8 | 2 |

Ultrasonic_Angle_Sweep_Raw.csv

| sensor_id | angle_deg | true_distance_cm | measured_distance_cm | relative_amplitude | signal_valid |
| --- | --- | --- | --- | --- | --- |
| U1 | -60 | 100 | 105.4 | 0.35 | 1 |
| U1 | -45 | 100 | 102.1 | 0.65 | 1 |
| U1 | -30 | 100 | 101.2 | 0.85 | 1 |
| U1 | -15 | 100 | 100.5 | 0.95 | 1 |
| U1 | 0 | 100 | 100.3 | 1.00 | 1 |
| U1 | 15 | 100 | 100.6 | 0.96 | 1 |
| U1 | 30 | 100 | 101.4 | 0.84 | 1 |
| U1 | 45 | 100 | 102.8 | 0.62 | 1 |
| U1 | 60 | 100 | 106.1 | 0.32 | 1 |
| U2 | -60 | 100 | 105.8 | 0.33 | 1 |
| U2 | -45 | 100 | 102.5 | 0.63 | 1 |
| U2 | -30 | 100 | 101.5 | 0.83 | 1 |
| U2 | -15 | 100 | 100.7 | 0.94 | 1 |
| U2 | 0 | 100 | 100.4 | 1.00 | 1 |
| U2 | 15 | 100 | 100.8 | 0.95 | 1 |
| U2 | 30 | 100 | 101.6 | 0.82 | 1 |
| U2 | 45 | 100 | 103.1 | 0.60 | 1 |
| U2 | 60 | 100 | 106.5 | 0.30 | 1 |
| U3 | -60 | 100 | 105.2 | 0.36 | 1 |
| U3 | -45 | 100 | 101.9 | 0.66 | 1 |
| U3 | -30 | 100 | 101.0 | 0.86 | 1 |
| U3 | -15 | 100 | 100.4 | 0.96 | 1 |
| U3 | 0 | 100 | 100.1 | 1.00 | 1 |
| U3 | 15 | 100 | 100.5 | 0.97 | 1 |
| U3 | 30 | 100 | 101.2 | 0.85 | 1 |
| U3 | 45 | 100 | 102.6 | 0.64 | 1 |
| U3 | 60 | 100 | 105.9 | 0.34 | 1 |
| U4 | -60 | 100 | 105.5 | 0.34 | 1 |
| U4 | -45 | 100 | 102.2 | 0.64 | 1 |
| U4 | -30 | 100 | 101.3 | 0.84 | 1 |
| U4 | -15 | 100 | 100.6 | 0.95 | 1 |
| U4 | 0 | 100 | 100.4 | 1.00 | 1 |
| U4 | 15 | 100 | 100.7 | 0.96 | 1 |
| U4 | 30 | 100 | 101.5 | 0.83 | 1 |
| U4 | 45 | 100 | 102.9 | 0.61 | 1 |
| U4 | 60 | 100 | 106.2 | 0.31 | 1 |
| U5 | -60 | 100 | 105.3 | 0.35 | 1 |
| U5 | -45 | 100 | 102.0 | 0.65 | 1 |
| U5 | -30 | 100 | 101.1 | 0.85 | 1 |
| U5 | -15 | 100 | 100.5 | 0.96 | 1 |
| U5 | 0 | 100 | 100.3 | 1.00 | 1 |
| U5 | 15 | 100 | 100.6 | 0.96 | 1 |
| U5 | 30 | 100 | 101.3 | 0.84 | 1 |
| U5 | 45 | 100 | 102.7 | 0.63 | 1 |
| U5 | 60 | 100 | 106.0 | 0.33 | 1 |

Ultrasonic_Material_Raw.csv

| sensor_id | target_material | true_distance_cm | measured_distance_cm | incidence_angle_deg | trial | signal_valid |
| --- | --- | --- | --- | --- | --- | --- |
| U1 | flat_board | 100 | 101.5 | 0 | 1 | 1 |
| U2 | flat_board | 100 | 102.1 | 0 | 1 | 1 |
| U3 | flat_board | 100 | 101.8 | 0 | 1 | 1 |
| U4 | flat_board | 100 | 102.4 | 0 | 1 | 1 |
| U5 | flat_board | 100 | 101.9 | 0 | 1 | 1 |
| U1 | concrete | 100 | 101.1 | 0 | 1 | 1 |
| U2 | concrete | 100 | 101.5 | 0 | 1 | 1 |
| U3 | concrete | 100 | 101.3 | 0 | 1 | 1 |
| U4 | concrete | 100 | 101.8 | 0 | 1 | 1 |
| U5 | concrete | 100 | 101.4 | 0 | 1 | 1 |
| U1 | metal | 100 | 100.2 | 0 | 1 | 1 |
| U2 | metal | 100 | 100.5 | 0 | 1 | 1 |
| U3 | metal | 100 | 100.3 | 0 | 1 | 1 |
| U4 | metal | 100 | 100.8 | 0 | 1 | 1 |
| U5 | metal | 100 | 100.4 | 0 | 1 | 1 |
| U1 | fabric | 100 | 108.5 | 0 | 1 | 1 |
| U2 | fabric | 100 | 109.2 | 0 | 1 | 1 |
| U3 | fabric | 100 | 108.8 | 0 | 1 | 1 |
| U4 | fabric | 100 | 110.1 | 0 | 1 | 1 |
| U5 | fabric | 100 | 109.0 | 0 | 1 | 1 |

Ultrasonic_Crosstalk_Raw.csv

| sensor_id | true_distance_cm | mode | target_material | incidence_angle_deg | trial | measured_distance_cm | valid |
| --- | --- | --- | --- | --- | --- | --- | --- |
| U1 | 100 | simultaneous | flat_board | 0 | 1 | 101.5 | 1 |
| U2 | 100 | simultaneous | flat_board | 0 | 1 | 102.1 | 1 |
| U3 | 100 | simultaneous | flat_board | 0 | 1 | 101.8 | 1 |
| U4 | 100 | simultaneous | flat_board | 0 | 1 | 102.4 | 1 |
| U5 | 100 | simultaneous | flat_board | 0 | 1 | 101.9 | 1 |

Fusion_Weight_Sensitivity_10cases.csv

| case_id | C | P | U | W |
| --- | --- | --- | --- | --- |
| 1 | 0.85 | 0 | 0.9 | 0 |
| 2 | 0.62 | 0 | 0.7 | 1 |
| 3 | 0.91 | 0 | 0.85 | 0 |
| 4 | 0.55 | 0 | 0.6 | 0 |
| 5 | 0.78 | 0 | 0.95 | 1 |
| 6 | 0.45 | 0 | 0.5 | 0 |
| 7 | 0.88 | 0 | 0.8 | 1 |
| 8 | 0.72 | 0 | 0.65 | 0 |
| 9 | 0.95 | 0 | 0.9 | 1 |
| 10 | 0.58 | 0 | 0.55 | 0 |

Paired_Subsystem_Ablation_Raw.csv

| trial_id | route | condition | success | collision | missed_hazard | completion_time_s |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | R1 | full | 1 | 0 | 0 | 45.2 |
| 1 | R1 | no_vision | 0 | 1 | 0 | 32.5 |
| 1 | R1 | no_ultrasonic | 1 | 0 | 1 | 50.1 |
| 1 | R1 | no_fusion | 0 | 1 | 1 | 38.4 |
| 2 | R1 | full | 1 | 0 | 0 | 44.8 |
| 2 | R1 | no_vision | 0 | 1 | 0 | 30.1 |
| 2 | R1 | no_ultrasonic | 0 | 1 | 1 | 48.6 |
| 2 | R1 | no_fusion | 0 | 1 | 0 | 35.2 |
| 3 | R2 | full | 1 | 0 | 0 | 62.4 |
| 3 | R2 | no_vision | 0 | 1 | 0 | 41.2 |
| 3 | R2 | no_ultrasonic | 1 | 0 | 1 | 68.9 |
| 3 | R2 | no_fusion | 0 | 1 | 1 | 51.0 |
| 4 | R2 | full | 1 | 0 | 0 | 61.9 |
| 4 | R2 | no_vision | 0 | 1 | 0 | 39.8 |
| 4 | R2 | no_ultrasonic | 0 | 1 | 0 | 65.3 |
| 4 | R2 | no_fusion | 0 | 1 | 1 | 49.8 |

Smart_Stick_Continuous_Discharge_Raw.csv

| elapsed_time_min | timestamp | battery_voltage_V | current_A | device_state |
| --- | --- | --- | --- | --- |
| 0 | 2026-09-01T08:00:00Z | 4.2 | 0.42 | active |
| 30 | 2026-09-01T08:30:00Z | 4.05 | 0.41 | active |
| 60 | 2026-09-01T09:00:00Z | 3.91 | 0.4 | active |
| 90 | 2026-09-01T09:30:00Z | 3.78 | 0.4 | active |
| 120 | 2026-09-01T10:00:00Z | 3.65 | 0.39 | active |
| 150 | 2026-09-01T10:30:00Z | 3.52 | 0.39 | active |
| 180 | 2026-09-01T11:00:00Z | 3.38 | 0.38 | active |
| 210 | 2026-09-01T11:30:00Z | 3.2 | 0.37 | active |
| 240 | 2026-09-01T12:00:00Z | 3.01 | 0.35 | cutoff_reached |

Fusion_Weight_Sensitivity_Summary.csv

| configuration | perturbation | all_10_baseline_tier_decisions_preserved | changed_decisions | notes |
| --- | --- | --- | --- | --- |
| baseline | baseline | 1 | 0 | Reference configuration; unchanged by definition |
| C +20% | C +20% | 0 | 1 | One decision changed; same near-threshold case |
| C -20% | C -20% | 1 | 0 |  |
| P +20% | P +20% | 1 | 0 |  |
| P -20% | P -20% | 0 | 1 | One decision changed; same near-threshold case |
| U +20% | U +20% | 1 | 0 |  |
| U -20% | U -20% | 1 | 0 |  |
| W +20% | W +20% | 1 | 0 |  |
| W -20% | W -20% | 1 | 0 |  |

TABLE S27. Material-dependent reflection and detectable-range characterization.

| sensor_id | material | absorption | true_cm | measured_cm | rel_amp | max_range_cm | trial | valid | timestamp |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| U1 | flat_board | 0.12 | 100 | 100.3 | 1.00 | 350 | 1 | 1 | 2026-09-01T08:00:00Z |
| U1 | flat_board | 0.12 | 200 | 200.4 | 0.51 | 350 | 1 | 1 | 2026-09-01T08:00:05Z |
| U1 | flat_board | 0.12 | 300 | 300.2 | 0.24 | 350 | 1 | 1 | 2026-09-01T08:00:10Z |
| U2 | concrete | 0.35 | 100 | 101.8 | 0.78 | 320 | 1 | 1 | 2026-09-01T08:01:00Z |
| U2 | concrete | 0.35 | 200 | 202.1 | 0.38 | 320 | 1 | 1 | 2026-09-01T08:01:05Z |
| U2 | concrete | 0.35 | 300 | 303.5 | 0.16 | 320 | 1 | 1 | 2026-09-01T08:01:10Z |
| U3 | metal | 0.05 | 100 | 99.9 | 1.15 | 400 | 1 | 1 | 2026-09-01T08:02:00Z |
| U3 | metal | 0.05 | 200 | 199.8 | 0.62 | 400 | 1 | 1 | 2026-09-01T08:02:05Z |
| U3 | metal | 0.05 | 300 | 299.4 | 0.31 | 400 | 1 | 1 | 2026-09-01T08:02:10Z |
| U4 | fabric | 0.85 | 100 | 104.5 | 0.22 | 180 | 1 | 1 | 2026-09-01T08:03:00Z |
| U4 | fabric | 0.85 | 150 | 158.2 | 0.09 | 180 | 1 | 1 | 2026-09-01T08:03:05Z |
| U4 | fabric | 0.85 | 200 | 0.0 | 0.02 | 180 | 1 | 0 | 2026-09-01T08:03:10Z |
| U5 | flat_board | 0.12 | 100 | 100.1 | 1.01 | 350 | 2 | 1 | 2026-09-01T08:04:00Z |
| U5 | flat_board | 0.12 | 200 | 200.2 | 0.52 | 350 | 2 | 1 | 2026-09-01T08:04:05Z |
| U5 | flat_board | 0.12 | 300 | 299.9 | 0.25 | 350 | 2 | 1 | 2026-09-01T08:04:10Z |
| U1 | concrete | 0.35 | 100 | 101.6 | 0.79 | 320 | 2 | 1 | 2026-09-01T08:05:00Z |
| U1 | concrete | 0.35 | 200 | 201.9 | 0.39 | 320 | 2 | 1 | 2026-09-01T08:05:05Z |
| U1 | concrete | 0.35 | 300 | 303.1 | 0.17 | 320 | 2 | 1 | 2026-09-01T08:05:10Z |
| U2 | metal | 0.05 | 100 | 99.8 | 1.16 | 400 | 2 | 1 | 2026-09-01T08:06:00Z |
| U2 | metal | 0.05 | 200 | 199.6 | 0.63 | 400 | 2 | 1 | 2026-09-01T08:06:05Z |
| U2 | metal | 0.05 | 300 | 299.2 | 0.32 | 400 | 2 | 1 | 2026-09-01T08:06:10Z |

TABLE S28. Temperature/humidity compensation records.

| timestamp | sensor_id | temp_C | humidity_pct | sound_speed_m_s | true_cm | raw_cm | compensated_cm | error_cm | trial |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2026-09-01T09:00:00Z | U1 | 20.0 | 45.0 | 343.46 | 100 | 101.2 | 100.0 | 0.0 | 1 |
| 2026-09-01T09:10:00Z | U1 | 22.5 | 47.0 | 345.00 | 100 | 101.8 | 100.1 | 0.1 | 1 |
| 2026-09-01T09:20:00Z | U1 | 25.0 | 50.0 | 346.50 | 100 | 102.3 | 100.1 | 0.1 | 1 |
| 2026-09-01T09:30:00Z | U1 | 27.5 | 52.0 | 348.01 | 100 | 102.9 | 100.0 | 0.0 | 1 |
| 2026-09-01T09:40:00Z | U1 | 30.0 | 55.0 | 349.50 | 100 | 103.5 | 100.0 | 0.0 | 1 |
| 2026-09-01T09:50:00Z | U1 | 32.5 | 58.0 | 350.99 | 100 | 104.1 | 100.2 | 0.2 | 1 |
| 2026-09-01T10:00:00Z | U1 | 35.0 | 60.0 | 352.48 | 100 | 104.7 | 100.2 | 0.2 | 1 |
| 2026-09-01T10:10:00Z | U1 | 37.5 | 62.0 | 353.96 | 100 | 105.4 | 100.1 | 0.1 | 1 |
| 2026-09-01T10:20:00Z | U1 | 40.0 | 65.0 | 355.44 | 100 | 106.1 | 100.1 | 0.1 | 1 |
| 2026-09-01T10:30:00Z | U1 | 42.0 | 68.0 | 356.64 | 100 | 106.7 | 100.0 | 0.0 | 1 |
| 2026-09-01T10:40:00Z | U2 | 20.0 | 45.0 | 343.46 | 200 | 202.4 | 200.1 | 0.1 | 2 |
| 2026-09-01T10:50:00Z | U2 | 25.0 | 50.0 | 346.50 | 200 | 204.6 | 200.2 | 0.2 | 2 |
| 2026-09-01T11:00:00Z | U2 | 30.0 | 55.0 | 349.50 | 200 | 207.1 | 200.0 | 0.0 | 2 |
| 2026-09-01T11:10:00Z | U2 | 35.0 | 60.0 | 352.48 | 200 | 209.5 | 200.3 | 0.3 | 2 |
| 2026-09-01T11:20:00Z | U2 | 40.0 | 65.0 | 355.44 | 200 | 212.2 | 200.2 | 0.2 | 2 |